# July 13, 2015 Successful interface to Toshiba Win 8.

| Windows settings: |  |
| --- | --- |
| In windows, use control panel to reach Ethernet Status | Control panel/network connections/ethernet/ properties |
| From Ethernet Status/properties disable IPV6. | Why: IPV6 active prevents IP address from staying in IPV4 |
| Set the IPV4 address to 192.168.1.2 set the subnet mask to: 255.255.255.0 | “OK” to close windows and store address. |
| CleverScope settings: |  |
| Use the Clevescope control panel to set up the connection. | CS control panel/settings/choose acquirer and set connection. |
| NIC adapter found should say: 192.168.1.2 | If it does not, you will need to troubleshoot. You may need to disable the wireless network adapter. The wireless network adapter can interfere with the NIC values in the cleverscope setup. You can enable the wireless adapter after the CS is working. |
| CleverScope IP settings: Under Cleverscope Acquirer Selected Type in the IP Address: 192.162.1.4 | The CS gets a different address from the PC. |
| Under Cleverscope Hardware configuration, Click on "Set IP Address" |  |
| After completion, it looks like: CleverScope Acquirer Selected: IP address = 192.168.1.4 NIC adapter found : 192.168.1.2 Subnet mask: 255.255.255.0 | Clever scope devices found on Local network Serial number = DK5649 (your serial number) Ip address = 192.168.1.4 Subnet mask = 255.255.255.0 gateway = 192.168.1.1 port = 53270 DHCP= "No, using static IP" Availability = "Not Available" time on = min, sec link test = <blank> |
| Now, click the CS control panel Auto button and traces should appear in the scope display. |  |

Figure 1: Control Panel

Figure 2:  Ethernet status

Figure 3: disable (TCP/IPv6)

Figure 4:  set IP address

Figure  5: Cleverscope Settings

Compare manual chapter 23 (page 146):

The only IP address ever referred to is: 192.168.1.2.

It does not explain that you need a different IP address in "CS acqurirer Selected/Ip address"

(The picture, however, does show 192.168.1.190, but this is not mentioned in the text)

Also , based on my experience, you need to disable IPV6.   If

IPV6 is enabled, the IPV4 will keep  hold the IP address.

In Cleverscope acquirer selected area, type in the fixed IP address you want to give the Cleverscope, and click

Set IP address. 192.168.1.20 in the example above.